Chapter I: Introduction to metrology

Introduction to Metrology:

Metrology word is derived from Greek words such as metro which means measurement and logy which means science. Metrology is the science of precision measurement. The engineer can say it is the science of measurement of lengths and angles and all related quantities like width, depth, diameter and straightness with high accuracy. Metrology demands pure knowledge of certain basic mathematical and physical principle.

Introduction to measurement:

Measurement is defined as the process of numerical evaluation of a dimension or the process of comparison with standard measuring instruments. The elements of measuring system include the instrumentation, calibration standard, environmental influence, human operator limitations and features of the work-piece.

Types of metrology:

- Legal metrology

Legal metrology is that part of metrology which treats units of measurements, methods of measurements and the measuring instruments, in relation to the technical and legal requirements. The activities of the service of legal metrology are,

1. Control of measuring instruments,

2. Testing of prototypes models of measuring instruments.

- Dynamic Metrology

Dynamic metrology is the technique of measuring small variations of a continuous nature. The technique has proved very valuable, and a record of continuous measurement, over a surface, for instance, has obvious advantages over individual measurements of an isolated character.

Methods of Measurements :

1. Direct method

2. Indirect method

3. Absolute of Fundamental method

4. Comparative method

5. Transposition method

6. Deflection method

7. Complementary method

8. Contact method

9. Contact less method of measurement are adopted depending upon the accuracy

1. Direct method of measurement:

This is a simple method of measurement, in which the value of the quantity to be measured is obtained directly without any calculations. For example, measurements by using scales, vernier callipers, micrometers, bevel protector etc. This method is not very accurate because it depends on human insensitiveness in making judgment.

2. Indirect method of measurement:

In indirect method the value of quantity to be measured is obtained by measuring other quantities which are functionally related to the required value. E.g. Angle measurement by sine bar, measurement of screw pitch diameter by three wire method etc.

3. Absolute or Fundamental method:

It is based on the measurement of the base quantities used to define the quantity.

For example, measuring a quantity directly in accordance with the definition of that quantity, or measuring a quantity indirectly by direct measurement of the quantities

Comparative method:

In this method the value of the quantity to be measured is compared with known value of the same quantity or other quantity practically related to it. So, in this method only the deviations from a master gauge are determined, e.g., dial indicators, or other comparators.

4. Transposition method:

It is a method of measurement by direct comparison in which the value of the quantity measured is first balanced by an initial known value A of the same quantity, and then the value of the quantity measured is put in place of this known value and is balanced again by another known value B. If the position of the element indicating equilibrium is the same in both cases, the value of the quantity to be measured is AB. For example, determination of amass by means of a balance and known weights, using the

Gauss double weighing.

5. Coincidence method:

It is a differential method of measurement in which a very small difference between the value of the quantity to be measured and the reference is determined by the observation of the coincidence of certain lines or signals. For example, measurement by vernier calliper micrometer.

6. Deflection method:

In this method the value of the quantity to be measured is directly indicated by a deflection of a pointer on a calibrated scale.

7. Complementary method:

In this method the value of the quantity to be measured is combined with a known value of the same quantity. The combination is so adjusted that the sum of these two values is equal to predetermined comparison value. For example, determination of the volume of a solid by liquid displacement.

8. Method of measurement by substitution:

It is a method of direct comparison in which the value of a quantity to be measured is replaced by a known value of the same quantity, so selected that the effects produced in the indicating device by these two values are the same.

9. Method of null measurement:

It is a method of differential measurement. In this method the difference between the value of the quantity to be measured and the known value of the same quantity with which it is compared is brought to zero.

GENERALIZED MEASUREMENT SYSTEM

A measuring system exists to provide information about the physical value of some variable being measured. In simple cases, the system can consist of only a single unit that gives an output reading or signal according to the magnitude of the unknown variable applied to it. However, in more complex measurement situations, a measuring system consists of several separate elements as shown in Figure1.1.

Table 1.1 Physical Quantities and its unit

Standards

The term standard is used to denote universally accepted specifications for devices. Components or processes which ensure conformity and interchangeability throughout a particular industry. A standard provides a reference for assigning a numerical value to a measured quantity. Each basic measurable quantity has associated with it an ultimate standard.

Working standards, those used in conjunction with the various measurement making instruments.

The national institute of standards and technology (NIST) formerly called

National Bureau of Standards (NBS), it was established by an act of congress in 1901, and the need for such body had been noted by the founders of the constitution

The following is the generalization of echelons of standards in the national measurement system.

1. Calibration standards

2. Metrology standards

3. National standards

1. Calibration standards: Working standards of industrial or governmental laboratories.

2. Metrology standards: Reference standards of industrial or Governmental laboratories.

3. National standards: It includes prototype and natural phenomenon of SI (Systems International), the world wide system of weight and measures standards.

Application of precise measurement has increased so much, that a single national laboratory to perform directly all the calibrations and standardization required by a large country with high technical development. It has led to the establishment of a considerable number of standardizing laboratories in industry and in various other areas. A standard provides a reference or datum for assigning a numerical value to a measured quantity.

Classification of Standards

To maintain accuracy and interchangeability it is necessary that Standards to be traceable to a single source, usually the National Standards of the country, which are further linked to International Standards. The accuracy of National Standards is transferred to working standards through a chain of intermediate standards in a manner given below.

• National Standards

• National Reference Standards

• Working Standards

• Plant Laboratory Reference Standards

• Plant Laboratory Working Standards

• Shop Floor Standards

Evidently, there is degradation of accuracy in passing from the defining standards to the shop floor standards. The accuracy of particular standard depends on a combination of the number of times it has been compared with a standard in a higher echelon, the frequency of such comparisons, the care with which it was done, and the stability of the particular standards itself.

Accuracy of Measurements

The purpose of measurement is to determine the true dimensions of a part. But no measurement can be made absolutely accurate. There is always some error. The amount of error depends upon the following factors:

• The accuracy and design of the measuring instrument

• The skill of the operator

• Method adopted for measurement

• Temperature variations

• Elastic deformation of the part or instrument etc.

Thus, the true dimension of the part cannot be determined but can only by approximate. The agreement of the measured value with the true value of the measured quantity is called accuracy. If the measurement of dimensions of a part approximates very closely to the true value of that dimension, it is said to be accurate. Thus the term accuracy denotes the closeness of the measured value with the true value. The difference between the measured value and the true value is the error of measurement. The lesser the error, more is the accuracy.

Precision

The terms precision and accuracy are used in connection with the performance of the instrument. Precision is the repeatability of the measuring process. It refers to the group of measurements for the same characteristics taken under identical conditions. It indicates to what extent the identically performed measurements agree with each other. If the instrument is not precise it will give different (widely varying) results for the same dimension when measured again and again. The set of observations will scatter about the mean. The scatter of these measurements is designated as σ, the standard deviation. It is used as an index of precision. The less the scattering more precise is the instrument.

Thus, lower, the value of σ, the more precise is the instrument.

Accuracy

Accuracy is the degree to which the measured value of the quality characteristic agrees with the true value. The difference between the true value and the measured value is known as error of measurement. It is practically difficult to measure exactly the true value and therefore a set of observations is made whose mean value is taken as the true value of the quality measured.

Distinction between Precision and Accuracy

Accuracy is very often confused with precision though much different. The distinction between the precision and accuracy will become clear by the following example.

Several measurements are made on a component by different types of instruments (A, B and C respectively) and the results are plotted. In any set of measurements, the individual measurements are scattered about the mean, and the precision signifies how well the various measurements performed by same instrument on the s

The difference between the mean of set the true value is called as error. Less the shows that the instrument A is precise since the results of number of measurements are close to the average value. However, there is a large difference (error) between the true value and the average value hence it is not accurate. The readings taken by the instruments are scattered much from the average value and hence it is not precise but accurate as there is a small difference between the average value and true value.

Factors affecting the accuracy of the Measuring System

The basic components of an accuracy evaluation are the five elements of a measuring system such as:

• Factors affecting the calibration standards.

• Factors affecting the work piece.

• Factors affecting the inherent characteristics of the instrument.

• Factors affecting the person, who carries out the measurements,

• Factors affecting the environment.

1. Factors affecting the Standard: It may be affected by:

-Coefficient of thermal expansion

-Calibration interval

-Stability with time

-Elastic properties

-Geometric compatibility

2. Factors affecting the Work piece: These are:

-Cleanliness

-Surface finish, waviness, scratch, surface defects etc.,

-Hidden geometry

-Elastic properties,-adequate datum on the work piece

-Arrangement of supporting work piece

-Thermal equalization etc.

3. Factors affecting the inherent characteristics of Instrument:

-Adequate amplification for accuracy objective

-Scale error

-Effect of friction, backlash, hysteresis, zero drift error

-Deformation in handling or use, when heavy work pieces are measured

-Calibration errors

-Mechanical parts (slides, guide ways or moving elements)

-Repeatability and readability

-Contact geometry for both work piece and standard.

4. Factors affecting person:

-Training, skill

-Sense of precision appreciation

-Ability to select measuring instruments and standards

-Sensible appreciation of measuring cost

-Attitude towards personal accuracy achievements

-Planning measurement techniques for minimum cost, consistent with precision requirements etc.

5. Factors affecting Environment:

-Temperature, humidity etc.

-Clean surrounding and minimum vibration enhance precision

-Adequate illumination

-Temperature equalization between standard, work piece, and instrument

-Thermal expansion effects due to heat radiation from lights

-Heating elements, sunlight and people

-Manual handling may also introduce thermal expansion.

Higher accuracy can be achieved only if, ail the sources of error due to the above five elements in the measuring system are analyzed and steps taken to eliminate them. The above analysis of five basic metrology elements can be composed into the acronym SWIPE, for convenient reference where, S – STANDARD W – WORKPIECE I – INSTRUMENT – P – PERSON E – ENVIRONMENT

SENSITIVITY

Sensitivity may be defined as the rate of displacement of the indicating device of an instrument, with respect to the measured quantity. In other words, sensitivity of an instrument is the ratio of the scale spacing to the scale division value. For example, if on a dial indicator, the scale spacing is 1.0 mm and the scale division value is 0.01 mm, then sensitivity is 100. It is also called as amplification factor or gearing ratio. If we now consider sensitivity over the full range of instrument reading with respect to measured quantities as shown in Figure the sensitivity at any value of y=dx/dy, where dx and dy are increments of x and y, taken over the full instrument scale, the sensitivity is the slope of the curve at any value of y.

The sensitivity may be constant or variable along the scale. In the first case we get linear transmission and in the second non-linear transmission. .

Sensitivity refers to the ability of measuring device to detect small differences in a quantity being measured. High sensitivity instruments may lead to drifts due to thermal or other effects, and indications may be less repeatable or less precise than that of the instrument of lower sensitivity.

Fig 1.2 Sensitivity of Measurement

Readability

Readability refers to the case with which the readings of a measuring Instrument can be read. It is the susceptibility of a measuring device to have its indications converted into meaningful number. Fine and widely spaced graduation lines ordinarily improve the readability. If the graduation lines are very finely spaced, the scale will be more readable by using the microscope; however, with the naked eye the readability will be poor. To make micrometers more readable they are provided with vernier scale. It can also be improved by using magnifying devices.

Calibration

The calibration of any measuring instrument is necessary to measure the quantity in terms of standard unit. It is the process of framing the scale of the instrument by applying some standardized signals. Calibration is a pre-measurement process, generally carried out by manufacturers. It is carried out by making adjustments such that the read out device produces zero output for zero measured input.

Repeatability

It is the ability of the measuring instrument to repeat the same results for the measurements for the same quantity, when the measurement are carried out-by the same observer,-with the same instrument,-under the same conditions,-without any change in location,-without change in the method of measurement-and the measurements are carried out in short intervals of time. It may be expressed quantitatively in terms of dispersion of the results.

Reproducibility

Reproducibility is the consistency of pattern of variation in measurement i.e. closeness of the agreement between the results of measurements of the same quantity, when individual measurements are carried out:

-by different observers

-by different methods

-using different instruments

-under different conditions, locations, times etc

ERRORS IN MEASUREMENTS

It is never possible to measure the true value of a dimension there is always some error. The error in measurement is the difference between the measured value and the true value of the measured dimension.

Error in measurement = Measured value - True value

The error in measurement may be expressed or evaluated either as an absolute error or as a relative error.

Absolute Error

True absolute error:

It is the algebraic difference between the result of measurement and the conventional true value of the quantity measured.

Apparent absolute error:

If the series of measurement are made then the algebraic difference between one of the results of measurement and the arithmetical mean is known as apparent absolute error. Relative Error:

It is the quotient of the absolute error and the value of comparison use or calculation of that absolute error. This value of comparison may be the true value, the conventional true value or the arithmetic mean for series of measurement. The accuracy of measurement, and hence the error depends upon so many factors, such as:

-calibration standard

-Work piece

-Instrument

-Person

-Environment etc.

Types of Errors

1. Systematic Error

These errors include calibration errors, error due to variation in the atmospheric condition Variation in contact pressure etc. If properly analyzed, these errors can be determined and reduced or even eliminated hence also called controllable errors. All other systematic errors can be controlled in magnitude and sense except personal error.

These errors results from irregular procedure that is consistent in action. These errors are repetitive in nature and are of constant and similar form.

2. Random Error

These errors are caused due to variation in position of setting standard and workpiece errors. Due to displacement of level joints of instruments, due to backlash and friction, these error are induced. Specific cause, magnitude and sense of these errors cannot be determined from the knowledge of measuring system or condition of measurement. These errors are non-consistent and hence the name random errors.

3. Environmental Error

These errors are caused due to effect of surrounding temperature, pressure and humidity on the measuring instrument. External factors like nuclear radiation, vibrations and magnetic field also leads to error. Temperature plays an important role where high precision is required. e.g. while using slip gauges, due to handling the slip gauges may acquire human body temperature, whereas the work is at 20°C. A 300 mm length will go in error by 5 microns which is quite a considerable error. To avoid errors of this kind, all metrology laboratories and standard rooms worldwide are maintained at 20°C.

4. Environmental Error

These errors are caused due to effect of surrounding temperature, pressure and humidity on the measuring instrument. External factors like nuclear radiation, vibrations and magnetic field also leads to error. Temperature plays an important role where high precision is required. e.g. while using slip gauges, due to handling the slip gauges may acquire human body temperature, whereas the work is at 20°C.

Calibration

Calibration is the process of checking the dimension and tolerances of a gauge, or the accuracy of a measurement instrument by comparing it to the instrument/gauge that has been certified as a standard of known accuracy. Calibration of an instrument is done over a period of time, which is decided depending upon the usage of the instrument or on the materials of the parts from which it is made. The dimensions and the tolerances of the instrument/gauge are checked so that we can come to whether the instrument can be used again by calibrating it or is it wear out or deteriorated above the limit value. If it is so then it is thrown out or it is scrapped. If the gauge or the instrument is frequently used, then it will require more maintenance and frequent calibration. Calibration of instrument is done prior to its use and afterwards to verify that it is within the tolerance limit or not.

Certification is given by making comparison between the instrument/gauge with the reference standard whose calibration is traceable to accepted National standard.

Chapter II - LINEAR AND ANGULAR MEASUREMENTS

LINEAR MEASURING INSTRUMENTS

Linear measurement applies to measurement of lengths, diameter, heights and thickness including external and internal measurements.. Linear measuring instruments are designed either for line measurements or end measurements. In end measuring instruments, the measurement is taken between two end surfaces as in micrometers, slip gauges etc.

The instruments used for linear measurements can be classified as:

1. Direct measuring instruments

2. Indirect measuring instruments

The Direct measuring instruments are of two types:

1. Graduated

2. Non Graduated

The graduated instruments include rules, vernier calipers, vernier height gauges, vernier depth gauges, micrometers, dial indicators etc.

The non graduated instruments include calipers, trammels, telescopic gauges, surface gauges, straight edges, wire gauges, screw pitch gauges, radius gauges, thickness gauges, slip gauges etc.

They can also be classified as

1. Non precision instruments such as steel rule, calipers etc.,

2. Precision measuring instruments, such as vernier instruments, micrometers, dial gauges etc.

SCALES

- The most common tool for crude measurements is the scale (also known as rules, or rulers).

- Although plastic, wood and other materials are used for common scales, precision scales use tempered steel alloys, with graduations scribed onto the surface.

- These are limited by the human eye. Basically they are used to compare two dimensions.

- The metric scales use decimal divisions, and the imperial scales use fractional divisions.

- Some scales only use the fine scale divisions at one end of the scale. It is advised that the end of the scale not be used for measurement. This is because as they become worn with use, the end of the scale will no longer be at a `zero' position.

- Instead the internal divisions of the scale should be used. Parallax error can be a factor when making measurements with a scale.

CALIPERS

Caliper is an instrument used for measuring distance between or over surfaces comparing dimensions of work pieces with such standards as plug gauges, graduated rules etc. Calipers may be difficult to use, and they require that the operator follow a few basic rules, do not force them, they will bend easily, and invalidate measurements made.

These instruments are very useful when dealing with hard to reach locations that normal measuring instruments cannot reach. Obviously the added step in the measurement will significantly decrease the accuracy.

VERNIER CALIPERS

The vernier instruments generally used in workshop and engineering metrology have comparatively low accuracy. The line of measurement of such instruments does not coincide with the line of scale. The accuracy therefore depends upon the straightness of the beam and the squareness of the sliding jaw with respect to the beam. To ensure the squareness, the sliding jaw must be clamped before taking the reading. The zero error must also be taken into consideration. Instruments are now available with a measuring range up to one meter with a scale value of 0.1 or 0.2 mm.

Types of Vernier Calipers

According to Indian Standard IS: 3651-1974, three types of vernier calipers have been specified to make external and internal measurements and are shown in figures respectively. All the three types are made with one scale on the front of the beam for direct reading.

Type A: Vernier has jaws on both sides for external and internal measurements and a blade for depth measurement